# ANNEX II + III:	TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

# Contract title: Medical equipment	p 1 /…

Publication reference: RORS00270/Secanj/TD4

Columns 1-2 should be completed by the contracting authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the contracting authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words ‘compliant’ or ‘yes’ are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item number | 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Holter 1+1 – 1set consisting of Holter of blood pressure 1pc Holter Device for continuous blood pressure measurement with a screen diagonal of at least 0.96 inches and a glass panel - Device dimensions: 106.5 x 62 x 27 mm ±5 mm The device has a self-locking structure - The memory stores up to 300 readings - Automatic re-measurement on failed attempt - Simple design with a maximum of one measurement button and one on/off button. - Static pressure range 0-300 mmHg ±3 mmHg or better - Resolution of pressure measurement 1 mmHg,pulse 1 bpm or better - Pulse measurement range 40-240 bpm - Cuff overinflation protection - Pressure measurement range: systolic - 25 to 290 mmHg, diastolic - 10 to 250 mmHg - Operating temperature from 5 to 40 ˚C - It has a C-type USB connection for connecting to a computer for programming measurements and data transfer - Minimum 40 characters available for entering the name of the hospital - Safety system - maximum inflation pressure limited to 300 mmHg, maximum measurement time to 120 s - Measurement duration minimum 24 h, optional in intervals: 5, 10, 15, 20, 30, 45, 60, 90, 120 min - Degree of protection IP22 or better - Continuous working mode - Pie Chart, Circadian and Statistical Analysis - Each table in the statistics includes the calculation of valid data, mean, standard deviation, maximum/minimum values, coefficient of variation, BP load, asleep dip, morning surge, morning pressure and AASI. - Battery powered - maximum two AA alkaline batteries - The device weighs a maximum of 170 g without battery and cuff Holter ECG - 1pc 12-channel ECG recorder - Standard software functions: to have dynamic learning and search function, histogram, atrial fibrillation analysis, heart rate variability analysis, Lorenz Plot, Waveform superposition heat map, ST analysis, QT analysis, page scan, pacemaker analysis - Optional: Waterfall analysis, Vector ECG functions, ventricular late potential analysis, T wave alternative analysis - Automatic analysis of ECG data - Duration of trend graphs displayed in horizontal coordinates – 6, 12 and 24 hours - Automatic removal of ECG signal interference - Automatic detection of arrhythmia - Grouping of QRS by type and by QRS maturity - Device dimensions: up to 92x56x20mm, weight up to 75 gr without battery - Degree of protection IP22 or better - Operating temperature 5-45 ˚C - Display: min 1.92″ color screen - The recorder has at least three buttons - The possibility of entering data (eg recording length, ID number) into the recorder without connecting to a computer - 24 hour recording capability, sampling rate 1024 Hz - Transfer of recordings from the recorder from an SD card or with a USB cable to a computer supported - The recorder must have an Event marker to indicate sudden changes - Marking of templates by color - color coded - The device is supplied with: ECG patient cable 10 leads, disposable ECG electrodes, software and key, case, SD card and reader, USB cable |  |  |  |
| 2 | ECG – 6pcs Six-channel electrocardiograph device - Minimum 7'' high resolution color LCD TFT touch display, minimum 800x480 - Displays the working status and ECG waveform of all 12 leads simultaneously - Print formats on a 6-channel printer; - 2*6+1(rhythm drain), 2*6, 3*4, 3*4+1(rhythm drain), - 4*3, 4*3+1(rhythm drain), 6*2, 6*2+1(rhythm drain), rhythm functions - Automatic analysis and interpretation - The capacity of the built-in memory (without expansion with memory cards) is a minimum of 1000 ECG records - Sensitivity setting: 2.5mm/mV, 5mm/mV, 10mm/mV, 20mm/mV, 40mm/mV, Auto - Built-in lithium battery with the ability to print a minimum of 190 ECG records - Thermal printer width min. 110 mm - Print speed setting: 5mm/s, 6.25mm/s, 10mm/s, 12.5mm/s, 25mm/s, 50mm/s, 100mm/s - Foil keyboard - Weight max. 2.3 kg - Ability to program the time stamp interval every 1/2/3/5/10/20/30/60 minutes |  |  |  |
| 3 | Portable Defibrillator -1pc Automated External Defibrillator (AED) intended for use in adult and pediatric patients - Ability to work in semi-automatic or fully automatic mode - Biphasic defibrillation waveform - Automatic ECG rhythm analysis with "Shock advised / No shock advised" decision making - Automatic energy charging when shock delivery is indicated - Defibrillation energy levels for adult patients: 150 J / 175 J / 200 J - Pediatric mode of operation with defined energy levels of 50 - 90 J - Voice instructions for the user during the entire resuscitation procedure - Visual indicators and control buttons for turning on the device, selecting the pediatric mode and delivering the shock - Device status indicator (Status LED) that shows the device's readiness for operation - Automatic periodic self-tests of the correctness of the device - Internal memory for storing AED interventions, including ECG signal and events, minimum 10 episodes of 15 hours of ECG and events - USB connection for data transfer - Weight device with electrodes and battery, maximum 2 kg - Battery capacity of more than 350 shocks at 200 J or a minimum of 36 hours of monitoring at temperatures of 25ºC - Automatically turn off the device after a period of inactivity to save battery - Built-in metronome for CPR feedback - Support for self-adhesive defibrillation electrodes for adult and pediatric patients - Portable design with integrated handle - Resistance to solid particles and liquids IP55 or better - Support for multiple voice guidance languages |  |  |  |
| 4 | Imuno-hemical analyser - 1pc Automated Electro-chemiluminescence Immunoassay Analyzer - Methodology: Electro-chemiluminescence Immunoassay (ECLIA) - Test mode: Batch, STAT - Throughput: min 86 tests/hour - Time to first result: min 9 mins - Reagent system: min 10 positions; 4-15 °C onboard refrigeration or better - Sample position: min 30 positions; Continuous loading - Sample volume and type: 5-100 µL; Serum and plasma or better - Incubation system: min 100 positions; 37±0.3 °C incubation or better - Probe system: Liquid level detection; Clot detection; Anti-collision function - Operation system: Windows 10, Touch screen, LIS compatible - Weight: max 100 kg |  |  |  |

Important Note:

- The Tenderer is obliged to submit a manufacturer's catalogue or other document in Serbian or English with the offer for the goods it offers, with a photo of the offered good and technical and other characteristics of the goods it offers, and from which the Contracting Authority can unequivocally determine that the offered goods correspond to the requirements requested in tender documentation.

- At latest prior delivery of equipment Contractor is obligated to submit ALIMS certificate (Decision of the Agency for Medicines and Medical Devices) for medical equipment in line with Serbian legislation and Arcticle 29 of Contract Special Conditions for this supply tender.