# ANNEX II + III:	TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

# Contract title: Supply of Lab equipment, chemical and consumables for analysis of honey	p 1 /…

Publication reference: RORS00240/RDASB/TD8

Columns 1-2 should be completed by the contracting authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the contracting authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words ‘compliant’ or ‘yes’ are not sufficient)

- Column 4 allows the tenderer to make comments onits proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item number | 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Liquid chromatograph -1 set Quaternary Pump - Low-pressure mixing principle, two pistons in series or equivalent high-precision pump design - Flow range 0.001–5 mL/min in 1 μL/min increments or better - Pump pressure range 2 – 60 MPa (20 – 600 bar or higher) for the entire flow range of 0.001–5 mL/min - Compressibility compensation Fully automated, independent of mobile phase composition - Flow accuracy ±1% or better - Flow precision≤ 0.07% RSD or better - Pressure pulsation <1% or <5 bar or better - Gradient pumping accuracy ±0.5% over the entire range or better - Precision gradient pumping <0.15% SD or better - Number of eluent lines 4 - Eluent degassing Integrated into pump, 4 channels minimum - Compatibility pH range 1–12.5 or better - Leak control Leak detectors Autosampler - Operating principle Split loop injection, flow-through-loop, flow-through-needle, or equivalent automatic injection principle - Pressure range 2–60MPa, (20–600 bar or higher) or better. - Injection volume 0.1–100 μL or higher - Injection volume accuracy ±1% or better - Injection volume precision <0.25% RSD or better - Injection linearity r ≥0.9999 or better - Injection cycle time 10 s or less - Carry over (UV) <0.003% (typical: <0.0009%) or better - Autosampler temperature range 4–40 ˚C or better - Autosampler capacity 150 vials or more (1.5 or 2 mL) - Leak control Leak detectors Column Heater - Temperature range minimum temperature: 5 ˚C or better (18 ˚C below ambient or better) - Maximum temperature: ≥85 ˚C - Temperature stability ±0.05 ˚C or better - Temperature accuracy ±0.5 ˚C or better - Column slots 2 Column slots up to 300mm or more - Temperature zones 1 or more - Leak control Leak detectors Refractive Index (RI) detector - Refractive Index range 1.00–1.75 RIU or better - Measurement range ±600 μRIU or better - Temperature control 30–55 °C or better - Linear range <5.0% RSD at 600 μRIU or better - Data acquisition rate Up to 50 Hz or better - Noise ≤2.5 nRIU or better - Drift ≤0.2 μRIU/hour or better - Flow cell Included in price - Safety options Leak detection. System Qualification - System qualification after installation with documentation (in accordance with 21 CFR Part 11 or equivalent) must be included and performed after system installation Chromatography software - Software for instrument control, sequence creation, data processing and report generation - The software should allow for later upgrade of the instrument with an MS or MS/MS spectrometer Computer and monitor - Brand name computer, i7 processor, 16GB ram, 1000GB HD or better - Monitor 23” or larger Note: - The price must include installation, qualification, and training at the user's site |  |  |  |
| 2 | Refractometer – 1pc - Range: Refractive index (nD) : at least 1.300000 - 1.715000 or wider; Brix measuring range : 0.00 - 100.00% - Resolution: Refractive index (nD) : 0.000001 or better; Brix: 0.01% or better; Temperature: 0.01°C or better - Measurement accuracy: ±0.000040 nD or better in the central measuring range; ±0.00010 nD or better for other ranges - Brix measurement accuracy: ±0.03% or better; ±0.05% or better for high Brix samples - Temperature control range: 5.00 - 75.00°C or wider - Installation, Calibration/validation and training must be included |  |  |  |
| 3 | UV/VIS spectrophotometer -1pc - Optical design – Dual beam, Czerny-Turner optical system - Wavelength measurement range – from 190 to 1100 nm; - Wavelength scanning speed - Selectable wavelength scanning speeds, with maximum scan speed 1600nm/min or better - Repeatability-< ± 0.2nm; - Spectral width: 2nm or narrower; - Resolution: 0.2 to 5 nm or equivalent - Wavelength accuracy: ± 0.5 nm - Sources: UV-Vis source covering 190-1100nm or better - Photometric linear range - from–2A to +3.0A or wider - Photometric accuracy ±0.002A at 0.5A and ±0.004A at 1A, or better; high-absorbance accuracy/linearity must be documented up to at least 2A - Photometric repeatability–±0.001Aat 1A or better - Light scattering: ≤1.0%T at 198 nm using KCl, ≤0.05%T at 220 nm using NaI, and ≤0.05%T at 340 nm or 370 nm using NaNO₂, or equivalent validated stray-light test according to manufacturer or pharmacopoeial conditionsNoise: - ≤0.00050A RMS or better under manufacturer-specified standard test conditions - Display: color touchscreenequivalent onboard/PC control interface - Languages: minimum English - Cuvette holders: supplied with automated multi-cell holder for at least 8 standard 10 mm cuvettes, or equivalent high-throughput cuvette changer The spectrophotometer must have the ability to upgrade for the following: - thermal printeror equivalent - Holder or changer for long-path cuvettes up to 50 mm and 100 mm, or equivalent long-path cell accessory - Peltier thermostated cuvette holder - Sipper accessory - Fiber optic dip probeor equivalent fiber-optic sampling accessory |  |  |  |
| 4 | FT IR spectrometer – 1pc - Spectral range: 7500–350 cm⁻¹ mid-infrared region or wider - Beamsplitter: KBr/Ge optimized for the mid-infrared region - Spectral resolution: 0.6 cm⁻¹ or better. - Signal-to-noise ratio: 35,000:1 (peak-to-peak, one minute), or better - Ordinate linearity (ASTM E1421 or equivalent): <0.15 %T deviation from 0.0 %T, or better. - Infrared source: long-life, high-stability MIR source, such as non-migrating hotspot single-point source or equivalent. - Laser: "solid-state", temperature-controlled diode laser with long working lifeor equivalent reference system with manufacturer-documented wavelength stability and lifetime. - Interferometer: dynamically aligned, permanently aligned, self-compensating, or equivalent high-stability interferometer design optimized for reproducible spectral throughput. - Detector: DTGS detector or equivalent - Modules: include in the offer a standard transmission module and an ATR module with a diamond crystal and attachments for liquid, powder samples. Controlled pressing of the sample is necessary. - User Interface/Connection: connection to laptop or computer - The software must support remote data access. - The operating software should support multicomponent search or equivalent and contamination analysis or equivalent, baseline correction, smoothing, Kramers–Kronig transformation or equivalent ATR correction, normalization, spectral mathematics, peak area and height analysis. - Chemometric software for quantitative analysis must be included with at least the Beer–Lambert law, PLS and CLS methodsor equivalent validated quantitative methods - It is mandatory to include spectrum basesor equivalent manufacturer-compatible libraries with comparable chemical coverage and total number of spectra: - Coatings Technology Library (2,507 spectra) - Hummel Polymer& Additives Library(2,011 spectra) - Polymers, Polymer Additives & Plasticizers Library (1,799 spectra) - Industrial Coatings Library (1,961 spectra) - Aldrich Condensed Phase Library (18,454 spectra) or equivalent - Aldrich Solvents Library (246 spectra) - Food Additives Library (519 spectra) - Sigma Biochemical Library (10,411 spectra) - Sigma Steroids Library (3,011 spectra) - The system must include an internal or supplied external certified NIST-traceable polystyrene standard for instrument qualification/verification - Diagnostics (minimum technical specifications): continuous or automated monitoring of key spectrometer components and/or environmental conditions, including temperature/humidity where applicable. Diagnostic status must be displayed via instrument indicator, software dashboard, or equivalent visual status display. - Deliver a PC, monitor, mouse, keyboard with the instrument. PC configuration: Intel Core Ultra 7 processor or equivalent/high-performance processor, 32 GB RAM or better - Hard Disk: minimum1TB SSD - Operating systems (minimum technical specifications): The software must be compatible with Microsoft Windows 11 - The offer should include training for chemometric software, working software |  |  |  |
| 5 | Microscope and camera – 1set 1. Microscope Optical system: - Infinity optical system - Trinocular head (30° tilted, 360° rotating) - Interpupillary distance: 50–75 mm - Dioptric correction on both eyepieces - Eyepieces: 10x / Ø 20 mm Objectives: - 4×, 10×, 20×, 40× spring-loaded, and 100× oil spring-loaded, with NA equal or better than 0.10, 0.25, 0.40, 0.65 and 1.25, respectively Lighting: - LED 3 W or equivalent - Continuous adjustment Condenser: - Centering, with diaphragm Mechanical stand: - 175 × 145 mm or larger - Displacement: at least 75 × 50 mm - Holder for two preparations 2. Camera Type and sensor: - 4K camera, 8 MP (3840 × 2160) - CMOS sensor, 1/1.8" or larger, or equivalent high-sensitivity sensor - Pixel: 2 × 2 µm or equivalent Performance: - 30 fps or higher - Exposure: 0.2 ms – 33.3 ms or wider Interface: - HDMI and USB 3.0or equivalent high-speed digital interface - C-mount Software: - manufacturer-supplied image acquisition and analysis software, supporting photography, video recording, calibration and measurements |  |  |  |
| 6 | Annealing furnace – 1pc - Tmax 1100 ° C - Heating on at least two sides with ceramic heating platesfor optimal temperature uniformity - Temperature uniformity of +/− 5 K with a closed fresh air damper in the empty working space according to DIN 17052-1 or equivalent, working temperatures above 800 ° C - Opening the door in a position to be able to use the working platform - Adjustable air inlet integrated in the door - Exhaust air outlet in the rear vent - Features of the controller: Controller with programmable temperature control, graphical or digital display, temperature programs, real-time/status display, fault/error memory, and programming in steps of 1 °C or 1 min or better. - Volume: 3 liters or more - Power: max 1.5 kW - Warm-up time (no load): max. 45 min. to Tmax according to manufacturer’s standard test method |  |  |  |
| 7 | Chemicals and standards – 1set - Mix of antibiotics and veterinary drugs relevant for the analysis of honey, with a minimum of 20 substances that comply with domestic and EU regulations. Can be a mix or individual standards. Suitable for LC-MS/MS analysis - Pesticide standards (can be mixed or individual standards (LC-MS/MS analysis): - ⁠pesticide mix that includes neonicotinoids - a minimum of 25 different common agricultural pesticides relevant to honey - ⁠mix varrocides (coumaphoe, tau-fluvalinate, amitraz + metabolites, flumethrin) - For the preparation of microscopic preparations (minimum 300 preparations) - ground microscope slides (76x26 mm), - coverslips (22x22 mm or larger) - glass sticks - Glycerin-gelatin (ready-made) - color carmine-acetate or fuchsin - immersion oil if x100 magnification is used - CRM: - ⁠Certified reference material, reference material, proficiency-test material, or matrix-matched fortified honey material containing one or more relevant pesticide, varroacide, or veterinary-drug residues - Sugar standards (at least analytical grade): - fructose - ⁠glucose - Sucrose - ⁠maltose - Standard HMF - Chemicals: - NaOH - ⁠phenolphthalein - ⁠CO₂-free deionized water prepared freshly according to the analytical method, or supplied water/reagents suitable for this purpose - The offer must include quantities sufficient for preparation, extraction, dilution, calibration and analysis of at least 300 samples of honey and bee products by LC-MS/MS, FTIR, UV-Vis, refractometry and HPLC-RI. - Minimum required solvents and reagents: - methanol, LC-MS grade - acetonitrile, LC-MS grade - ethyl acetate, analytical or LC-MS grade - n-hexane or hexanes, analytical grade - isopropanol, analytical or LC-MS grade - formic acid, LC-MS grade - acetic acid - ammonium formate, LC-MS grade - ammonium acetate, LC-MS grade - hydrochloric acid, concentrated - ethanol 95% - any additional extraction salts, buffers, filtration materials, vials, caps, inserts, or equivalent consumables required by the offered analytical workflow |  |  |  |
| 8 | Freeze dryer – 1pc - Electric power consumption: up to 2kW - Drying capacity: 1 - 5kg or more - Number of trays: min 4 - Minimal temperature of condenser: -40oC or lower |  |  |  |

Important Notes:

- The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

- The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

- Unless otherwise specified, the requirements in these Technical Specifications are presented as a minimum standard which the offered goods must meet.

- For all required standards required in technical specification tenderer may offer equipment with equivalent standard compliance if applicable.

Warranty:

Tenderers must provide local reliable warranty service agent providing maintenance and the rapid supply of equipment spare parts and consumables for the Warranty duration of one year.

Offer must include warranty service description including:

•	Service organisation contact data including name, postal address, telephone number, fax number and e-mail address;

•	Help Desk (phone) support, which must be available during working hours, 8AM – 4PM;

•	Guaranteed maximum response time to submitted maintenance support request (fax or e-mail) of 1 (one) working day;

•	Guaranteed that any requests for services will be attended to within 24 hours;

•	Guarantee that all items can be repaired or alternatively replaced within a maximum of 72 hours;